## NXentry
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
experiment_title | Title of the experiment | NXentry:title | NX_CHAR |  | F | 0 | 0
experiment_location | City and country where ex. Took place | NXentry:experiment_location | NX_CHAR |  | F | 0 | 0
experiment_start_date | Start date and time of ex. in ISO 8601 format, ideally including UTC offset. | NXentry:experiment_start_date | NX _DATE_TIME |  | F | 0 | 0
experiment_end_date | Start date and time of ex. in ISO 8601 format, ideally including UTC offset. | NXentry:experiment_end_date | NX_DATE_TIME |  | F | 0 | 0
experiment_summary | Descriptive summary of ex. | NXentry:experiment_description | NX_CHAR |  | T | 0 | 0
experiment_documentation | Documents describing the experiment | NXentry:experiment_documentation |  | NXnote | T | 0 | 0
experiment_identifier | Unique identifier for the experiment. | NXentry:experiment_identifier | NX_CHAR |  | T | 0 | 0
experiment_run_cycle | Some facilities organize their beamtime into run cycles. | NXentry:run_cycle | NX_CHAR |  | T | 0 | 0
experiment_institution | Name of the institution hosting the facility | NXentry:experiment_institution | NX_CHAR |  | F | 0 | 0
experiment_facility | Name of the experimental facility | NXentry:experiment_facility | NX_CHAR |  | F | 0 | 0
experiment_laboratory | Name of the laboratory or beamline | NXentry:experiment_laboratory | NX_CHAR |  | F | 0 | 0
title | Title of the entry, referring to a specific measurement. | NXentry:title | NX_CHAR |  | T | 0 | 0
entry_identifier | Identifier of the measurement | NXentry:entry_identifier | NX_CHAR |  | T | 0 | 0
start_time | Start date and time of ex. in ISO 8601 format, ideally including UTC offset. | NXentry:start_time | NX_DATE_TIME |  | T | 0 | 0
end_time | Ending date of the measurement | NXentry:end_time | NX_DATE_TIME |  | T | 0 | 0
duration | Duration of the measurement | NXentry:duration | NX_FLOAT:NX_TIME |  | T | 0 | 0
collection_time | Time spent actually collecting data | NXentry:collection_time | NX_FLOAT:NX_TIME |  | T | 0 | 0
definition | Does the file correspond to a standardized structure? | NXentry:definition | NX_CHAR |  | T | 0 | 0

## NXuser
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
name | Name of user responsible for this entry | NXentry:NXuser:name | NX_CHAR |  | T | 0 | 0
role | Role of user (local contact, principal investigator, etc.) | NXentry:NXuser:role | NX_CHAR |  | T | 0 | 0
affiliation | Affiliation of user | NXentry:NXuser:affiliation | NX_CHAR |  | T | 0 | 0
address | Address of user | NXentry:NXuser:address | NX_CHAR |  | T | 0 | 0
telephone_number | Telephone number of user | NXentry:NXuser:telephone_number | NX_CHAR |  | T | 0 | 0
email | E-mail address of user | NXentry:NXuser:email | NX_CHAR |  | T | 0 | 0
facility_user_id | User identifier defined by the facility | NXentry:NXuser:facility_user_id | NX_CHAR |  | T | 0 | 0

## NXinstrument
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
instrument_name | Name or model of the equipment | NXentry:NXinstrument:name | NX_CHAR |  | T | 0 | 0
attribute | Acronym or other shorthand name | NXentry:NXinstrument:name:\@short_name | NX_CHAR |  | T | 0 | 0
description | Short description of the instrument, possibly reference to a technical publication | NXentry:NXinstrument:description | NX_CHAR |  | F | 0 | 0

## NXsource
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
name | Name of the source | NXentry:NXinstrument:NXsource:name | NX_CHAR |  | T | 0 | 0
type | Type of radiation source | NXentry:NXinstrument:NXsource:type | NX_CHAR |  | T | 0 | 0
probe | Type of radiation used as probe | NXentry:NXinstrument:NXsource:probe | NX_CHAR |  | T | 0 | 0
photon_energy | The center photon energy of the source, before it is monochromatized or converted | NXentry:NXinstrument:NXsource:photon_energy | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
center_wavelength | The center wavelength of the source, before it is monochromatized or converted | NXentry:NXinstrument:NXsource:center_wavelength | NX_FLOAT:NX_WAVELENGTH |  | F | 0 | 0
pulse_energy | For pulsed sources, the energy of a single pulse | NXentry:NXinstrument:NXsource:pulse_energy | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
peak_power | For pulsed sources, the pulse energy divided by the pulse duration | NXentry:NXinstrument:NXsource:peak_power | NX_FLOAT:NX_POWER |  | F | 0 | 0
power | Average source power | NXentry:NXinstrument:NXsource:power | NX_FLOAT:NX_POWER |  | T | 0 | 0
emittance_x | Horizontal emittance | NXentry:NXinstrument:NXsource:emittance_x | NX_FLOAT:NX_EMITTANCE |  | T | 0 | 0
emittance_y | Vertical emittance | NXentry:NXinstrument:NXsource:emittance_y | NX_FLOAT:NX_ EMITTANCE |  | T | 0 | 0
sigma_x | horizontal beam size | NXentry:NXinstrument:NXsource:sigma_x | NX_FLOAT:NX_LENGTH |  | T | 0 | 0
sigma_y | vertical beam size | NXentry:NXinstrument:NXsource:sigma_y | NX_FLOAT:NX_LENGTH |  | T | 0 | 0
flux | Photon flux of the source | NXentry:NXinstrument:NXsource:flux | NX_FLOAT:NX_FLUX |  | T | 0 | 0
energy | Particle beam energy/excitation voltage | NXentry:NXinstrument:NXsource:energy | NX_FLOAT:NX_ENERGY |  | T | 0 | 0
current | Storage ring current/X-ray tube current | NXentry:NXinstrument:NXsource:current | NX_FLOAT:NX_CURRENT |  | T | 0 | 0
frequency | Pulsed source frequency (repetition rate) | NXentry:NXinstrument:NXsource:frequency | NX_FLOAT:NX_FREQUENCY |  | T | 0 | 0
pulse_width | Duration of the source pulses | NXentry:NXinstrument:NXsource:pulse_width | NX_FLOAT:NX_TIME |  | T | 0 | 0
distribution | Energy/wavelength spectrum of the source | NXentry:NXinstrument:NXsource:distribution |  | NXdata | T | 2 | [1,nx],[1,nx]
number_of_bunches | Number of bunches in the storage ring | NXentry:NXinstrument:NXsource:number_of_bunches | NX_INT |  | T | 0 | 0
bunch_length | Temporal length of the bunch | NXentry:NXinstrument:NXsource:bunch_length | NX_FLOAT:NX_TIME |  | T | 0 | 0
bunch_distance | Time between bunches | NXentry:NXinstrument:NXsource:bunch_distance | NX_FLOAT:NX_TIME |  | T | 0 | 0
mode | Source operating mode (single or multi bunch) | NXentry:NXinstrument:NXsource:mode | NX_CHAR |  | T | 0 | 0
top_up | Is the storage ring in top up? | NXentry:NXinstrument:NXsource:top_up | NX_BOOLEAN |  | T | 0 | 0
bunch_pattern | Plottable description of the bunch pattern | NXentry:NXinstrument:NXsource:bunch_pattern |  | NXdata | T | 2 | [1,nx],[1,nx]
rms_jitter | For pump-probe experiments, the synchronization error between pump and probe. | NXentry:NXinstrument:NXsource:rms_jitter | NX_FLOAT:NX_TIME |  | T | 0 | 0
bunch_number_start | Some facilities tag each bunch. First bunch of the measurement | NXentry:NXinstrument:NXsource:bunch_number_start | NX_INT |  | F | 0 | 0
bunch_number_end | Last bunch of the measurement | NXentry:NXinstrument:NXsource:bunch_number_end | NX_INT |  | F | 0 | 0

## NXbeam
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
distance | Distance from the sample | NXentry:NXinstrument:NXbeam:distance | NX_FLOAT:NX_LENGTH |  | T | 0 | 0
pulse_energy | Energy of a single pulse at the diagnostic point | NXentry:NXinstrument:NXbeam:pulse_energy | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
average_power | Average power at the diagnostic point | NXentry:NXinstrument:NXbeam:average_power | NX_FLOAT:NX_POWER |  | F | 0 | 0
incident_energy | The center photon energy at the diagnostic point | NXentry:NXinstrument:NXbeam:incident_energy | NX_FLOAT:NX_ENERGY |  | T | 0 | 0
incident_energy_spread | FWHM energy linewidth of the incident photons. | NXentry:NXinstrument:NXbeam:incident_energy_spread | NX_NUMBER:NX_ENERGY |  | F | 0 | 0
incident_wavelength | The center wavelength energy at the diagnostic point | NXentry:NXinstrument:NXbeam:incident_wavelength | NX_FLOAT:NX_WAVELENGTH |  | T | 0 | 0
incident_wavelength_spread | FWHM wavelength linewidth of the incident photons. | NXentry:NXinstrument:NXbeam:incident_wavelength_spread | NX_NUMBER:NX_WAVELENGTH |  | T | 0 | 0
incident_polarization | Incident polarization as a Stokes vector | NXentry:NXinstrument:NXbeam:incident_polarization | NX_FLOAT:NX_ANGLE |  | F | 1 | [1,4]
extent | beam size vector | NXentry:NXinstrument:NXbeam:extent | NX_FLOAT:NX_LENGTH |  | T | 1 | [1,2]
fluence | Incident fluence at the sample | NXentry:NXinstrument:NXbeam:fluence | NX_FLOAT:NX_ANY |  | F | 0 | 0
pulse_duration | Duration of the pulses at the diagnostic point | NXentry:NXinstrument:NXbeam:pulse_duration | NX_FLOAT:NX_TIME |  | F | 0 | 0
chirp_spectrum | Spectrum of pulse frequency vs time. | NXentry:NXinstrument:NXbeam:chirp_spectrum |  | NXdata | F | 2 | [1,nx],[1,nx]
chirp_type | The type of chirp implemented | NXentry:NXinstrument:NXbeam:chirp_type | NX_CHAR |  | F | 0 | 0
chirp_GDD | Group delay dispersion of the pulse for linear chirp | NXentry:NXinstrument:NXbeam:chirp_GDD | NX_FLOAT:NX_TIME |  | F | 0 | 0
flux | flux incident on beam plane area | NXentry:NXinstrument:NXbeam:flux | NX_FLOAT:NX_FLUX |  | T | 0 | 0

## NXmonochromator
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
energy | Central energy | NXentry:NXinstrument:NXmonochromator:energy | NX_FLOAT:NX_ENERGY |  | T | 0 | 0
energy_error | Energy standard deviation | NXentry:NXinstrument:NXmonochromator:energy_error | NX_CHAR |  | T | 0 | 0
distribution | Spectrum or live diagnostic | NXentry:NXinstrument:NXmonochromator:distribution |  | NXdata | T | 2 | [1,nx],[1,nx]

## NXmanipulator
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
name | Name | NXentry:NXinstrument:NXmanipulator:name | NX_CHAR |  | F | 0 | 0
description | Description | NXentry:NXinstrument:NXmanipulator:description | NX_CHAR |  | F | 0 | 0
type | Type of manipulator, Hexapod, Rod, etc. | NXentry:NXinstrument:NXmanipulator:type | NX_CHAR |  | F | 0 | 0
coordinates | Coordinates of the manipulator position (x,y,z) | NXentry:NXinstrument:NXmanipulator:coordinates | NX_FLOAT:NX_LENGTH |  | F | 1 | [1,ncoord]
angles | Angles of the manipulator orientation (polar,tilt,azimuth) | NXentry:NXinstrument:NXmanipulator:angles | NX_FLOAT:NX_ANGLE |  | F | 1 | [1,nangles]
coordinate_values | Effective positions assumed by the manipulator during the measurement. | NXentry:NXinstrument:NXmanipulator:coordinate_values | NX_FLOAT:NX_LENGTH |  | F | 2 | [ncoord,npos]
angle_values | Effective angles assumed by the manipulator during the measurement. | NXentry:NXinstrument:NXmanipulator:angle_values | NX_FLOAT:NX_ANGLE |  | F | 2 | [nangles,npos]
cryocoolant | Is cryocoolant flowing through the manipulator? | NXentry:NXinstrument:NXmanipulator:cryocoolant | NX_BOOL |  | F | 0 | 0
cryostat_temperature | Temperature of the cryostat (coldest point) | NXentry:NXinstrument:NXmanipulator:cryostat_temperature | NX_FLOAT:NX_TEMPERATURE |  | F | 0 | 0
heater_power | Power in the heater for temperature control. | NXentry:NXinstrument:NXmanipulator:heater_power | NX_FLOAT:NX_POWER |  | F | 0 | 0
sample_temperature | Temperature at the closest point to the sample | NXentry:NXinstrument:NXmanipulator:sample_temperature | NX_FLOAT:NX_TEMPERATURE |  | F | 0 | 0
drain_current | Current to neutralize the photoemission current | NXentry:NXinstrument:NXmanipulator:drain_current | NX_FLOAT:NX_CURRENT |  | F | 0 | 0
sample_bias | Possible bias of the sample with trespect to analyser ground | NXentry:NXinstrument:NXmanipulator:sample_bias | NX_FLOAT:NX_CURRENT |  | F | 0 | 0
POSITIONER | class to describe the motors that are used in the manipulator | NXentry:NXinstrument:NXmanipulator:POSITIONER |  | NXpositioner | F | 0 | 0
TRANSFORMATIONS | class to describe the transformations used. | NXentry:NXinstrument:NXmanipulator:TRANSFORMATIONS |  | NXtransformations | F | 0 | 0

## NXelectronanalyser
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
description | Free text description of the type of the detector | NXentry:NXinstrument:NXelectronanalyser:description | NX_CHAR |  | F | 0 | 0
name | Name or model of the equipment | NXentry:NXinstrument:NXelectronanalyser:name | NX_CHAR |  | F | 0 | 0
attribute | Acronym or other shorthand name | NXentry:NXinstrument:NXelectronanalyser:name:\@short_name | NX_CHAR |  | F | 0 | 0
energy_resolution | Overall energy resolution (FWHM of gaussian broadening) | NXentry:NXinstrument:NXelectronanalyser:energy_resolution | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
momentum_resolution | Overall momentum resolution (FWHM) | NXentry:NXinstrument:NXelectronanalyser:momentum_resolution | NX_FLOAT:NX_WAVEVECTOR |  | F | 0 | 0
spatial_resolution | Overall spatial resolution (Airy disk radius) | NXentry:NXinstrument:NXelectronanalyser:spatial_resolution | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
energy_scan_mode | fixed or sweep | NXentry:NXinstrument:NXelectronanalyser:acquisition_mode | NX_CHAR |  | F | 0 | 0
fast_axes | List of the axes that are acquired symultaneously by the detector | NXentry:NXinstrument:NXelectronanalyser:fast_axes | NX_CHAR |  | F | 1 | [1,nfa]
slow_axes | List of the axes that are acquired, in order of decreasing speed. For example energy in a swept scan or spin. | NXentry:NXinstrument:NXelectronanalyser:slow_axes | NX_CHAR |  | F | 1 | [1,nsa]
COLLECTION_COLUMN | class to describe the electron collection (spatial and momentum imaging) column | NXentry:NXinstrument:NXelectronanalyser:COLLECTION_COLUMN |  | NXcollectioncolumn | F | 0 | 0
ENERGY_DISPERSION | class to describe the energy dispersion section | NXentry:NXinstrument:NXelectronanalyser:ENERGY_DISPERSION |  | NXenergydispersion | F | 0 | 0
SPIN_DISPERSION | class to describe the spin dispersion section | NXentry:NXinstrument:NXelectronanalyser:SPIN_DISPERSION |  | NXspindispersion | F | 0 | 0
DETECTOR | class to describe the electron detector | NXentry:NXinstrument:NXelectronanalyser:DETECTOR |  | NXdetector | F | 0 | 0

## NXcollectioncolumn
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
scheme | Scheme of the electron collection lens, i.e. standard, deflector, PEEM, momentum microscope, etc. | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:scheme | NX_CHAR |  | F | 0 | 0
extractor_voltage | Voltage applied to the extractor lens | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:extractor_voltage | NX_FLOAT:NX_VOLTAGE |  | F | 0 | 0
extractor_current | Indicating leakage, field emission or arc currents to the extractor lens | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:extractor_current | NX_FLOAT:NX_CURRENT |  | F | 0 | 0
working_distance | Distance between sample and detector entrance | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:working_distance | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
mode | Labelling of a standard lens setting | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:mode | NX_CHAR |  | T | 0 | 0
lens_names | Set of names of electron optic lenses | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:lens_names | NX_CHAR |  | F | 1 | [1,nl]
lens_voltages | Array of corresponding voltages | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:lens_voltages | NX_FLOAT:NX_VOLTAGE |  | F | 1 | [1,nl]
projection | The space projected in the angularly dispersive directions, real or reciprocal | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:projection | NX_CHAR |  | F | 0 | 0
magnification | The magnification of the projected image in the angularly dispersive direction | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:magnification | NX_FLOAT:NX_DIMENSIONLESS |  | F | 0 | 0
field_aperture | The size and position of the field aperture inserted in the column | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:field_aperture |  | NXaperture | F | 0 | 0
contrast_aperture | The size and postion of the contrast aperture inserted in the column. | NXentry:NXinstrument:NXelectronanalyser:NXcollectionlens:contrast_aperture |  | NXaperture | F | 0 | 0
DEFLECTOR | deflectors in the collection column section | NXentry:NXinstrument:NXelectronanalyser:NXcollectioncolumn:DEFLECTOR |  | NXdeflector | F | 0 | 0
LENS | Individual lenses in the collection column section | NXentry:NXinstrument:NXelectronanalyser:NXcollectioncolumn:LENS |  | NXlens | F | 0 | 0

## NXenergydispersion
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
scheme | tof, hemispherical, cylindrical, mirror, retarding grid, etc. | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:scheme | NX_CHAR |  | F | 0 | 0
pass_energy | Pass energy in dispersive electrodes | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:pass_energy | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
center_energy | Center of the energy window | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:center_energy | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
entrance_slit | Size, position and shape of the entrance slits in dispersive analyzers | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:entrance_slit |  | NXaperture | F | 0 | 0
exit_slit | Size, position and shape of the exit slits in dispersive analyzer | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:exit_slit |  | NXaperture | F | 0 | 0
diameter | Diameter of the dispersive orbit | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:diameter | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
energy_scan_mode | fixed or sweep | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:energy_scan_mode | NX_CHAR |  | F | 0 | 0
tof_distance | Length of the tof drift electrode | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:tof_distance | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
DEFLECTOR | deflectors in the energy dispersive section | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:DEFLECTOR |  | NXdeflector | F | 0 | 0
LENS | Individual lenses in the energy dispersive section | NXentry:NXinstrument:NXelectronanalyser:NXenergydispersion:LENS |  | NXlens | F | 0 | 0

## NXspindispersion
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
spin_filter_type | Type of spin detector, VLEED, SPLEED, Mott, etc. | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_filter_type | NX_CHAR |  | F | 0 | 0
spin_FoM | Figure of merit of the spin detector | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_filter_FoM | NX_FLOAT:NX_DIMENSIONLESS |  | F | 0 | 0
spin_shermann_function | Effective Shermann function, calibrated spin selectivity factor | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_filter_shermann_function | NX_FLOAT:NX_DIMENSIONLESS |  | F | 0 | 0
spin_scattering_energy | Energy of the spin-selective scattering | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_scattering_energy | NX_FLOAT:NX_ENERGY |  | F | 0 | 0
spin_scattering_angle | Angle of the spin-selective scattering | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_scattering_angle | NX_FLOAT:NX_ANGLE |  | F | 0 | 0
spin_target | Name of the target | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_target | NX_CHAR |  | F | 0 | 0
spin_target_preparation | Preparation procedure of the spin target | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_target_preparation | NX_CHAR |  | F | 0 | 0
spin_target_preparation_date | Date of last preparation of the spin target | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:spin_target_preparation_date | NX_DATE_TIME |  | F | 0 | 0
DEFLECTOR | deflectors in the spin dispersive section | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:DEFLECTOR |  | NXdeflector | F | 0 | 0
LENS | Individual lenses in the spin dispersive section | NXentry:NXinstrument:NXelectronanalyser:NXspindispersion:LENS |  | NXlens | F | 0 | 0

## NXdetector
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
amplifier_type | Type of electron amplifier, MCP, channeltron, etc. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:amplifier_type | NX_CHAR |  | F | 0 | 0
detector_type | Description of the detector type, DLD, Phosphor+CCD, CMOS. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:detector_type | NX_CHAR |  | F | 0 | 0
detector_voltage | Voltage applied to detector. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:detector_voltage | NX_FLOAT:NX_VOLTAGE |  | F | 0 | 0
amplifier_voltage | Voltage applied to the amplifier. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:amplifier_voltage | NX_FLOAT:NX_VOLTAGE |  | F | 0 | 0
amplifier_bias | The low voltage of the amplifier migh not be the ground. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:amplifier_bias | NX_FLOAT:NX_VOLTAGE |  | F | 0 | 0
sensor_size | Size of each imaging sensor chip on the detector. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:sensor_size | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
sensor_count | Number of imaging sensor chips on the detector. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:sensor_count | NX_INT:NX_UNITLESS |  | F | 0 | 0
sensor_pixel_size | Physical size of the pixels of the imaging chip on the detector. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:sensor_pixel_size | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
sensor_pixels | Number of raw active elements in each dimension. Important for swept scans. | NXentry:NXinstrument:NXelectronanaliser:NXdetector:sensor_pixels | NX_INT:NX_UNITLESS |  | F | 0 | 0
DEFLECTOR | deflectors in the detector section | NXentry:NXinstrument:NXelectronanalyser:NXdetector:DEFLECTOR |  | NXdeflector | F | 0 | 0
LENS | Individual lenses in the detector section | NXentry:NXinstrument:NXelectronanalyser:NXdetector:LENS |  | NXlens | F | 0 | 0
DATA | raw data output from the detector | NXentry:NXinstrument:NXelectronanalyser:NXdetector:DATA |  | NXdata | F | 0 | 0

## NXsample
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
name | Simple and descriptive name of the sample | NXentry:NXsample:name | NX_CHAR |  | T | 0 | 0
sample_id | Identification number or signatures of the sample used. | NXentry:NXsample:sample_id | NX_CHAR |  | F | 0 | 0
state | Physical state of the sample | NXentry:NXsample: state | NX_CHAR |  | F | 0 | 0
purity | Chemical purity of the sample | NXentry:NXsample: purity | NX_FLOAT:NX_UNITLESS |  | F | 0 | 0
surface_orientation | Surface termination of the sample (if crystalline) | NXentry:NXsample: orientation | NX_CHAR |  | F | 0 | 0
layer | Number of layers of the sample (e.g. bulk, monolayer, pentalayer, etc.) | NXentry:NXsample: layer | NX_CHAR |  | F | 0 | 0
space_group | Space group of the sample (if crystalline) | NXentry:NXsample: space_group | NX_CHAR |  | T | 0 | 0
chemical_formula | Chemical formula of the sample | NXentry:NXsample: chemical_formula | NX_CHAR |  | T | 0 | 0
chemical_name | Full chemical name of the sample | NXentry:NXsample: chemical_name | NX_CHAR |  | F | 0 | 0
chem_id_cas | CAS registry number of the sample chemical content. | NXentry:NXsample: chem_id_cas | NX_CHAR |  | F | 0 | 0
temperature | Temperature of the sample during the measurement. | NXentry:NXsample:sample_temperature | NX_FLOAT:NX_TEMPERATURE |  | T | 0 | 0
drain_current | Current to neutralize the photoemission current | NXentry:NXsample:drain_current | NX_FLOAT:NX_CURRENT |  | T | 0 | 0
pressure | Pressure of gas surrounding the sample at time of measurement. | NXentry:NXsample: pressure | NX_FLOAT:NX_PRESSURE |  | T | 0 | 0
gas | Gases might be fluxed on the surface for various reasons. Chemical designation, or residual. | NXentry:NXsample: gas | NX_CHAR |  | F | 0 | 0
electric_field | Electric field applied across the sample. Different from sample_bias. | NXentry:NXsample:electric_field | NX_FLOAT:NX_VOLTAGE |  | T | 0 | 0
attribute | Vector determining the direction of the electric field | NXentry:NXsample:electric_field:\@direction | NX_FLOAT:NX_VOLTAGE |  | T | 1 | [1,3]
magnetic_field | Magnetic field applied to the sample during the measurement. | NXentry:NXsample:magnetic_field | NX_FLOAT:NX_ANY |  | T | 0 | 0
attribute | Vector determining the direction of the magnetic field | NXentry:NXsample:magnetic_field:\@direction | NX_FLOAT:NX_ANY |  | T | 1 | [1,3]
stress_field | Stress field applied to the sample during the measurement. | NXentry:NXsample:stress_field | NX_FLOAT:NX_PRESSURE |  | T | 0 | 0
attribute | Vector determining the direction of the stress field | NXentry:NXsample:magnetic_field:\@direction | NX_FLOAT:NX_PRESSURE |  | T | 1 | [1,3]
thickness | Sample thickness | NXentry:NXsample:thickness | NX_FLOAT:NX_LENGTH |  | T | 0 | 0
surface_dopant | Element of evaporated surface dopant such as alkali or other | NXentry:NXsample:surface_dopant | NX_CHAR |  | F | 0 | 0
surface_dopant_coverage | Nominal thickness of the evaporated dopant | NXentry:NXsample:surface_dopant_coverage | NX_FLOAT:NX_LENGTH |  | F | 0 | 0
bias | Voltage applied to sample and sample holder. | NXentry:NXsample:bias | NX_FLOAT:NX_VOLTAGE |  | F | 0 | 0
drain_current | Drain current due to photoemission | NXentry:NXsample:drain_current | NX_FLOAT:NX_CURRENT |  | F | 0 | 0
growth_method | Sample growth method (e. g. molecular beam epitaxy, chemical vapor deposition etc.) | NXentry:NXsample:growth_method | NX_FLOAT:NX_CURRENT |  | F | 0 | 0
preparation_method | E.g. UHV cleave, sputtering and annealing, etc. | NXentry:NXsample:preparation_method | NX_FLOAT:NX_CURRENT |  | F | 0 | 0
preparation_date | Date of sample preparation | NXentry:NXsample:preparation_date | NX_DATE_TIME |  | T | 0 | 0
vendor | Name of the sample vendor (company or research group) | NXentry:NXsample:vendor | NX_CHAR |  | F | 0 | 0
substrate_material | Material of the substrate in direct contact with the sample. | NXentry:NXsample:substrate_material | NX_CHAR |  | F | 0 | 0
substrate_state | Physical state of the substrate, similar options to sample_state | NXentry:NXsample:substrate_state | NX_CHAR |  | F | 0 | 0
substrate_vendor | Vendor of the substrate material | NXentry:NXsample:substrate_material | NX_CHAR |  | F | 0 | 0
sample_notes | Further notes. | NXentry:NXsample:notes |  | NXnote | F | 0 | 0

## NXprocess
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
program_name | Name of the processing program or script | NXentry:NXprocess:program_name | NX_CHAR |  | F | 0 | 0
program_version | Version of the program or script | NXentry:NXprocess:program_version | NX_CHAR |  | F | 0 | 0
program_codebase | Set of libraries/tools used in the program/script | NXentry:NXprocess:program_codebase | NX_CHAR |  | F | 0 | 0
program_sequence | Sequence of transformations applied to the data, as some are non-commutative | NXentry:NXprocess:program_sequence | NX_CHAR |  | F | 0 | 0
REGISTRATION | Class to describe the operations of image registration | NXentry:NXprocess:REGISTRATION |  | NXregistration | F | 0 | 0
DISTORTION | class to describe the operations of image distrortion correction | NXentry:NXprocess:DISTORTION |  | NXdistortion | F | 0 | 0
CALIBRATION | class to describe the operations of calibration | NXentry:NXprocess:CALIBRATION |  | NXcalibration | F | 0 | 0

## NXdistortion
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
applied | Has the distortion correction been applied? | NXentry:NXprocess:NXdistortion:applied | NX_BOOL |  | F | 0 | 0
symmetry | Number of elementary symmetry operations. | NXentry:NXprocess:NXdistortion:symmetry | NX_INT:NX_UNITLESS |  | F | 0 | 0
symmetry_angle | Degrees of rotation for a symmetry operation. | NXentry:NXprocess:NXdistortion:symmetry_angle | NX_FLOAT:NX_ANGLE |  | F | 0 | 0
original_centre | Coordinates of the symmetry centre point. | NXentry:NXprocess:NXdistortion:original_centre | NX_FLOAT:NX_UNITLESS |  | F | 1 | [1,2]
original_points | Coordinates of the relevant symmetry points. | NXentry:NXprocess:NXdistortion:original_points | NX_FLOAT:NX_UNITLESS |  | F | 1 | [2,nsym]
Field | 2D matrix mapping the original distorted field in the undistorted one. | NXentry:NXprocess:NXdistortion:field | NX_FLOAT:NX_UNITLESS |  | F | 1 | [ndx,ndy]

## NXregistration
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
applied | Has the registration been applied? | NXentry:NXprocess:NXregistration:applied | NX_BOOL |  | F | 0 | 0
new_centre | Coordinates of the new symmetry centre point. | NXentry:NXprocess:NXregistration:new_centre | NX_FLOAT:NX_UNITLESS |  | F | 1 | [1,2]
rotation_centre | Coordinates of the rotation centre. | NXentry:NXprocess:NXregistration:rotation_centre | NX_FLOAT:NX_UNITLESS |  | F | 1 | [1,2]
scaling | Scaling factor to compensate shrinking from distortion correction. | NXentry:NXprocess:NXregistration:scaling | NX_FLOAT:NX_DIMENSIONLESS |  | F | 1 | [1,2]
TRANSFORMATION | To describe the operations of image registration (combinations of rigid translations and rotations) | NXentry:NXprocess:NXregistration:TRANSFORMATION |  | NXtransformation | F | 0 | 0

## NXcalibration
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
applied | Has the calibration been applied? | NXentry:NXprocess:NXcalibration:applied | NX_BOOL |  | F | 0 | 0
coefficients | Array of coefficients of the calibration function for non linear calibration | NXentry:NXprocess:NXcalibration:coefficients | NX_FLOAT:NX_ANY |  | F | 1 | [1,ncoeff]
features | Array of position of features along the axis to be calibrated | NXentry:NXprocess:NXcalibration:features | NX_FLOAT:NX_ANY |  | F | 1 | [1,nfeat]
destinations | Array of destinations in the calibrated axis | NXentry:NXprocess:NXcalibration:destinations | NX_FLOAT:NX_ANY |  | F | 1 | [1,nfeat]
files | Files used for the calibration | NXentry:NXprocess:NXcalibration:files | NX_CHAR |  | F | 1 | [1,nfeat]
scaling | Scaling parameter for linear calibration | NXentry:NXprocess:NXcalibration:scaling | NX_FLOAT:NX_ANY |  | F | 0 | 0
offset | Offset parameter for linear calibration | NXentry:NXprocess:NXcalibration:offset | NX_FLOAT:NX_ANY |  | F | 0 | 0
calibrated_axis | A vector representing the axis after calibration, matching the data length | NXentry:NXprocess:NXcalibration:calibrated_axis | NX_FLOAT:NX_ANY |  | F | 1 | [1,ncal]
conversion_table | 2 vectors, representing how the raw points get mapped on the calibrated axis. | NXentry:NXprocess:NXcalibration:conversion_table | NX_FLOAT:NX_ANY |  | F | 2 | [1,ncal],[1,ncal]

## NXaperture
Name | Documentation | Nexus hierarchy | type:units | NX class | Exists | Rank | Dim
description | Description of the type of aperture, i.e. slit, pinhole, iris | NXentry:NXinstrument:NXaperture:description | NX_CHAR |  | T | 0 | 0
shape | Shape of the aperture, i.e. straight or curved for slits, circle, square hexagon, octagon, bladed for pinholes and irises. | NXentry:NXinstrument:NXaperture:shape | NX_CHAR |  | F | 0 | 0
size | The relevant dimension for the aperture, i.e. slit width, pinhole and iris diameter | NXentry:NXinstrument:NXaperture:size | NX_NUMBER:NX_LENGTH |  | F | 0 | 0