## Sheet1
NionSwift, AutoSTEM2 (AS2), every variable can only be a number, SI base units, Chris (@Nion)
JSON term	What is it ?	Units
high_tension	total acceleration voltage extraction + acceleration	V
defocus		m	where measured ?
lens currents in A
camera image or scanned image have different sets of metadata
eht	isA high_tension	a.u.
PMTBF_gain	voltage on dynodes pmt	a.u.
PMTDF_gain	two different photomultipliers, PMT2_gain HAADF	a.u.	array of dynodes
StageOutA	tilt	rad
StageOutB		rad
StageOutX	position (eucentric) stage, mechanical position, optical axis	m
StageOutY		m
StageOutZ		m
AS2 as coordinate Andreas Mittelberger Nion software allows take optical image orient stage coordinates on the image, set up a optical microscope, digital camea optical microscope image
C10	polynomial coefficients for the geometric aberrations in the plane	m
C12.a	a and b relation to x,y and via can be related via coordinates corrector coordinates
C12.b	Krivanek
C21.a
C21.b
C23.a
C23.b
C30
…
C50
C1 const power	two counter winding, change ratio of the two currents	A
C2 const power	scale Amps real amps related to 100 kV scaling factor	A
C3 const power		A
AS2 lens excitations future
PMT2_gain	HAADF PMT?	a.u.
SuperFEG emission current	if two high, broaden your probe	A
G_2Db	gun deflector
LastTuneCurrent	??
probe_ha	probe half angle (semi convergence angle)	rad
HAADF_inner_ha	haadf annulus angular range inner and out	rad
HAADF_outer_ha		rad
GeometricProbeSize	diameter, calculated from high tension and aberrations	m

Scan	how the scan coils were
SuperScan	product name, scan box driving the scan coils
rotation
rotation_deg
Idea: make own scan box, attach
scan_context_size		px
scan_size	identical to scan_context_size are identical so no sub-scan fov_nm/scan_size left and right boundaries	px
scan_device_parameters
flyback_time
calibration style	space, or angular
external_clock_mode	external scan function generator

scan_device_properties
ac_line_sync	50Hz wait before you move to the next line, sync scanning with, 0, 1	boolean
AS2 is autonomous scan box is hierarchical not below
flyback_time_us	move from two extremal positions, scaning compressing pixels, gets sync, ac_line_sync
pixel_time_us	time spend at a pixel
	512+n, line sync
requested_pixel_time_us	pixel time, scan box, 6 MHz accuracy
rotation_deg	the scan is rotated	deg
rotation_rad	probe is not rectangular, STEM you rotate the scan	rad
magBoard tool relevant for calibration your magnitude nxcollection? digital analog converter settings, uint16s !

hardware source
exposure	total scan time ? Npix*Npix * pixel_time_us = exposure	s
frame_index	??
channel_id	HADF is channel a, each channel one pmt, so a,b,c possible
channel_name	HADF
channel_index	[0, 1, 2] → [a,b,c]
ScanBox talks to AS2 and Swift equal rank
BeamCurrent	remains to be implemented

Normal HADF image mode